Durham Split-Pair Horizontal magnet.

The Superconducting Magnet

Central field					:	15 tesla at 4.2 K

Central field homogeneity			:	< 0.4% over a 10 mm DSV

< 5.2% over a 40 mm DSV

Field profiles					:	see appendix 2

Operating current				:	116 A nominal at 15 tesla

Inductance					:	244 H nominal

Stored energy					:	1.6 Mjoule nominal

Maximum field sweep rate			:	0.2 tesla min-1

Field stability in the persistent mode		:	< 1x 10-4 relative hr-1

Vertical access to the high field region		:	40 mm diameter

## The Cryostat

Liquid helium capacity				:	55 L

Liquid helium autonomy			:	> 5 days with magnet in persistent mode

at 4.2 K and zero heat load from experiment or variable temperature insert in static mode

Liquid Nitrogen capacity			:	58 L

Liquid nitrogen autonomy			:	> 6 days

Sample access					:	NW 40 (via neck baffle adapter)

Overall system dimensions (see drawing)	:	Max diameter – 781 mm

Max height – 1840 mm

Sample entry to magnetic centre – 1320 mm

Windows			:	4 off horizontal optical access, two parallel to field,

two perpendicular to field. Cryostat outer tail fitted

with 38mm diameter Spec. B quartz windows.

Cryostat re-entrant bore tubes fitted with 8mm diameter Spec.

B quartz windows.

Optical Access					:	F10 nominal normal and parallel to field direction

Optical path in liquid helium			:	< 10mm radial

Optical Tail variable temperature insert

Temperature range				:	< 2 to 300 K

Temperature stability				:	< ( 0.1 K

Operating mode				:	Dynamic with sample in liquid / gas flow

Sample region diameter			:	10 mm

Sample access entry port			:	NW25

Temperature control sensor			:	Calibrated Cernox

Windows			:	4 off horizontal optical access, two parallel to field,

two perpendicular to field, fitted with Spec. B quartz

windows. VTI outer tail fitted with 8mm diameter

Spec. B quartz windows. VTI sample space fitted with 5mm diameter Spec. B quartz windows.

Modulation coil

Maximum modulation amplitude		:	( 0.005 tesla

Maximum modulation frequency		:	50 Hz

Operating current				:	(2.5 amp nominal

Modulating field homogeneity			:	13.2% over a 10 mm DSV

ECSM electronic control console including:

ITC503 digital temperature controller:

Single channel input

Upgradeable to 3 channel input

4½ digit display

IEEE488 and RS232 interfaces

Standard calibration curves built-in

Auto-ranging

Auto-PID setting

'ObjectBench' software for data-logging and range handling

Ability to control auto needle valves

ILM211 nitrogen and helium double level meter and probes:

Sensing method: capacitance for nitrogen; superconducting wire for helium

Nitrogen probe - maximum length 2 m, diameter 6.35 mm

Helium probe maximum length 1 m, diameter 4.8 or 9.5 mm

Variable sample rate for helium level (5 s to 33 hr)

Low level fill indicators for auto-fill systems

Audible low level alarm

Digital display

Display units: percent

Display resolution 0.1%

RS232 interface

Includes all cables and rack mounting

EPS40 helium pumping system comprising:

40 m3/hr rotary pump

Base pressure 1 mbar

2 m long flexible pumping line and valve

## ‘T’ Piece, valve and pressure gauge for Lambda pumping port

Electrical Power Requirements:- 240V / 50Hz